LMView

Multidevice Universal

Help & User Manual

G-Tek Corporation Pvt. Ltd.

Vadodara, Gujarat, India

www.gtek-india.com

Copyright © 2026 G-Tek Corporation Pvt. Ltd. All rights reserved.

Index

About G-Tek	4

What’s New in LMView	5

LMView terms	6

Mathematical Calculations	7

System Requirements	8

Application Window	9

Application Window	10

Installation	13

Uninstallation	19

Overview	20

Starting LMView MD	21

Connect To LittleMaster	22

Go Offline	24

Setup Batch on Multiple Device	25

Download Batch from LM	37

Calibrate LittleMaster	43

Stop Running Batch	45

Current Data	46

Open Downloaded Batches	49

Filter Table Data	53

Format Tabular Data	55

Calculate MKT	58

Unit Conversion	59

View Graph	60

Zoom In Graph	61

Zoom Out Graph	62

Save Graph	63

Print Graph	64

Transfer Axis	65

Select Channels	66

Graph Properties	67

Export to CSV	74

PDF Report	76

Error Description	84

# About G-Tek

LMView

With the CR-2010 and SR-2000 series of Circular and Strip Chart Recorder G-Tek has won the confidence of the customer for reliability, performance and responsive services.

The new Little Master LM series once again proves the capability and tenacity of G-Tek Corporation to bring forth world-class products at affordable price. This LM Reusable series of small data loggers record the temperature and %RH at the set interval of time. Wide temperature range, Polynomial based linearization, High accuracy, customized LCD display, Long Battery Life and low power consumption are some of the salient features. LmTransit is a one time use data logger Available in variant memory size with prefixed store interval. This logger comes pre-programmed from factory and user just needs pull the tab to start the recording. The associated LMView software helps download the stored data and review them. User can easily download the software from the website www.gtek-india.com For any query or suggestion please get in touch with us. Before installing and operating the instrument, please take time to go through the manual, which will enable you to get the most from your Little Master.

Our new series owes much to the feedback received from our customer and we shall always welcome your suggestions and comments on any subject of our product.

G-Tek Corporation Pvt. Ltd.

Vadodara

# What’s New in LMView

The enhanced layout tools, feature improvements, redesigned user interface, and improved workflow can boost your work efficiency and give you a more enjoyable work experience.

Updated user interface

Attempts have been made to introduce new icons, menus, and controls within the application to create a fresh look and provide a more intuitive working environment. Graphics and reporting capabilities have been greatly improved.

Tag into device

Apply tag into device for capture any particular data or time for future use.

Multidevice connection

Now user can connect multiple devices of same type of device family at a time.

Multidevice batch configuration

User can configure multiple devices batch at a time.

Multi Download

Multiple connected devices batch data can be download parallelly.

# LMView terms

Before you get started with LMView, please familiarize yourself with the terms being used.

| Term | Description |
| --- | --- |
| LM | LittleMaster device like LM Temperature only, LM Temperature and %RH etc. |
| MKT | Mean Kinetic Temperature |
| Temperature (°C) | Temperature in Degree Celsius |
| Temperature (°F) | Temperature in Degree Fahrenheit |
| Temperature (°K) | Temperature in Degree Kelvin |
| Humidity (%RH) | It is a ratio that compares the amount of water vapor in the air with the amount of water vapor that would be present in the air at saturation. |
| Dew Point (°F/°K) | The dew point temperature is the temperature at which the air can no longer hold all of its water vapor, and some of the water vapor must condense into liquid water. At 100% relative humidity, the dew point temperature and real temperature are the same, and clouds or fog can begin to form. The unit of dew point is same as that of temperature unit. |
| RS232-C | Serial Communication. |
| USB | Universal Serial Bus. |
| OnLine | When LM is physically connected to PC. User can perform various activities like configuring the batch or downloading the stored data. User can also see the current data as measured by LM |
| OffLine | When LM is physically not connected to the PC. User can perform activities like viewing, reporting and exporting the previously stored data. |

# Mathematical Calculations

Dew Point:

The dew point calculation uses the recorded values of relative humidity – RH (%) and temperature – T (°C) and the following formula:

Where: K=273.14°C; B=5420°K ; T is the temperature at that point in Deg C; RH is the measured %RH

Mean Kinetic Temperature:

It is a simplified way of expressing the overall effect of temperature fluctuations during storage or transit of perishable goods. In other words, MKT is a calculated, single temperature that is analogous to the effects of temperature variations over a period. The formula for Mean Kinetic Temperature is as follows:

Where:

is the mean kinetic temperature in kelvin
  is the activation energy (typically within 60–100 kJ·mol-1 for solids or liquids)
  is the gas constant
  to  are the temperatures at each of the sample points in kelvins
   is the number of temperature sample points

# System Requirements

Operating Systems: Windows 10; Windows 11;

Processor: i5 1GHz processor or equivalent (Minimum); i5 2GHz Pentium processor or equivalent (Recommended)

RAM: 8 GB (Minimum);

Hard Disk: Up to 5 GB of available space may be required

CD or DVD Drive: Not required

Display: 800 x 600, 256 colors (Minimum); 1024 x 768 high color, 32-bit (Recommended)

# Application Window

When you launch LMView, the application window opens containing a main window. Although more than one main window can be opened, you can apply commands to the active main window only.

The LMView application window is shown below.

| Part | Description |
| --- | --- |
| 1. Title bar | The area displaying the title of the application name and mode. |
| 2. Menu bar | The area containing pull-down menu options |
| 3. Toolbar | A detachable bar that contains shortcuts to menu and other commands |
| 4. Working Panel | The panel is called a working panel which will convert into different panels as per the current activity e.g. Device Identification Panel, Device Connection Panel, TableData Panel, Current Data Panel etc. |

# Application Window

Application commands are accessible through the menu bar, toolbars, and toolbox.

Main Toolbar

| Image | Command Name | Command Work |
| --- | --- | --- |
|  | Open Downloaded Batch | Open the previously Downloaded Batch and display the selected batch data in tabular form. |
|  | View Graph | Open the graphical view of tabular data being displayed |
|  | Filter Tabular Data | Use this button to filter displayed tabular data giving new start time and stop time. |
|  | Export to CSV | Exports the displayed data as comma separated values which can be easily opened using an Excel like spreadsheet application. |
|  | Export to PDF | Export the displayed data in PDF format to the specified file. |
|  | Connect to LM | Open the connection dialog to communicate to LM. |
|  | Disconnect from LM | While in Online mode use this button of the toolbar to disconnect the LM and to continue work in offline mode. This can also be used to connect another device after disconnecting from first device. |
|  | Current Data | Shows current data in current data panel. |
|  | Setup Batch on Multiple Device | Open the configuration dialog for configuration of alarm and set up (configure) the batch on multiple devices. |
|  | Stop Running Batch | Stops data storage in LM. |
|  | Download stored batch from LM | Opens the save dialog to Download and save stored data from the LM to the PC. |
|  | Calculate MKT | Calculates Mean Kinetic Temperature for the current tabular data. |
|  | Unit Conversion | Converts temperature from one unit to another. |
|  | Help | Help file for quick reference. |
|  | Exit | Closes the application. |

Graph Toolbar

| Image | Command Name | Command Work |
| --- | --- | --- |
|  | Save | Save Graph as PNG, JPG file. |
|  | Print | Print Graph using installed printer. |
|  | Transfer Axis | Transfer a channel's Y axis from Y to Y2 or Y2 to Y. |
|  | Select Channels | Displays selected channels on graph. |
|  | Graph Properties | Opens the Graph Properties dialog box. |
|  | Close Window | Exit graph view. |
|  | Zoom In | Zoom both the axis. |
|  | Zoom-inX | Zoom X axis. |
|  | Zoom-inY | Zoom Y axis. |
|  | Zoom Out | Zoom Out both the axis. |
|  | Zoom-outX | Zoom Out X axis. |
|  | Zoom-outY | Zoom Out Y axis. |

# Installation

This section describes the interactive installation guide that will take you through the installation process.

it starts initializing which may take some time.

After initialization process is completed, it starts to install application. In welcome dialog, click "Next" to proceed further.

Please read the license agreement. If you agree to the terms, then check on the accept terms. If you disagree, you have an option to click "cancel" and exit the setup program.

Once you accept the terms click on "Next" to proceed further.

The software requires the installation of .NET Framework 8.0. If the software does not find it, it will display the message. If you chose "Yes" then it starts to install ".Net Framework 8.0". If you chose "No" then it skips this process and continues with application installation readme process. If .Net Farmwork 8.0 is already installed into machine then it does not display this message and continues with application installation process.

After that Microsoft .Net Framework 8.0 setup wizard opens, click "Download it now" to proceed further.

Accept the "End-User License Agreement" and click "Install" to start the installation process.

After completion of all setup process, click on "Finish" to continue with installation of LMView MD.

Please read application information and click "Next" to proceed.

The default directory where the application will be installed is displayed. If you want to change the location, click on "Browse" and point to the directory where you want to install. It is recommended to install at the default location. Once the directory is selected, click on "Next" to proceed further.

The default path where the program will be resident, is displayed. This is the path you will see in your start menu. It is recommended to keep the default path and proceed further by clicking on "Next".

The required disk space and available disk space is displayed. Please ensure that you have sufficient disk space to install the program. Click on "Next" to continue.

Review the installation summary and once satisfied, click on "Install" to start installation process.

The progress bar indicates the installation process.

After successful installation it will show the following message. Click on finish button to end the installation process.

# Uninstallation

To Uninstall the application, you have two options.

(1) from Start  Go to control panel and remove the application from there.

Click on Start  Control Panel  Add or Remove Programs.

In Add or Remove Programs select LMView MD and click "Uninstall" for uninstallation process.

On click on Uninstall following message will appear to confirm that you want to uninstall the application or not, if yes click on Yes button.

This dialog displays entire uninstallation process and after successful uninstallation click "Close" to exit.

# Overview

The LMView MD software is designed to program the various parameters for data recording of LM*. It also helps download the stored data to PC and to view, analyse, print and export the stored data.

User can configure various batch parameters, download stored batch data, view data in tabular and graphical formats and can also view the current data. User can also see the minimum, maximum and average values for a batch in tabular format. User can also filter the tabular data giving new start and stop time. It is possible to switch between various temperature units in table. Dew point is calculated from temperature and %RH data. The software can calculate MKT (mean kinetic temperature) for specified activation energy. User can view the data in graphical format on click of a button. Zoom in and Zoom out facility is provided in graphical format. User can change various parameters like title, axis colours, axis scales etc. A special password protected calibration menu is provided. This helps the user to calibrate the LM on site. LM can be corrected for both scale and offset error.

An export to CSV format is provided to view or process the data in more generic spreadsheet editor like Excel or it can be imported to any other user defined database. A utility to convert the temperatures from one unit to another unit is also provided

The most common tasks and commands are available as buttons on the main toolbar.

# Starting LMView MD

- To start LMViewE051, click on the icon on desktop.

- Application will always start in online mode.

- Once LM is started, it shows the main window with two panels – Left and right.

- Left panel have a list of devices to connect into application and a dropdown to select other device family and used to display device information.

- Right panel display RTC and Current Data of device.

There are two modes of LMView – Online and Offline

Online: When LM is physically connected to PC. User can perform various activities like configuring the batch or downloading the stored data. User can also see the current data as measured by LM.

Offline: When LM is physically not connected to the PC. User can perform activities like viewing, reporting and exporting the previously stored data.

# Connect To LittleMaster

Connect Multiple LM to the free USB ports of the PC.

LMView always starts in an online mode. Connect the devices to any of the available USB ports of the PC.

- Click on Refresh Devices on Left Panel Select the devices Click on Connect Selected.

OR Click on the Icon to open the connection dialog box to connect the device.

- Click on offline Connect to Little Master.

- LMView application will try to connect with LM device.

- If application got connected with device, then it will show you the current data of device with dropdown to switch between connected devices as shown in the below image.

On the right-hand side, it shows Current Data of device with dropdown above to switch between connected devices and on left hand side it will show the Product code of the device, Serial No and also of the LM and the firmware Version of the LM device.

# Go Offline

"Go Offline" command is enabled when LMView is in online mode.

Click on offline Go Offline

OR Click on this Icon to get into offline mode of the application.

When offline command is applied, Connected LM devices get disconnected and LmView application gets into the offline mode, User will get to see the following screen in the offline mode. To communicate with LM again, user need to refresh devices and then select devices from any connected device family and then click on connect selected to go "Online" (Connect to LM) again. In offline mode user can open downloaded batch for viewing into tabular form.

# Setup Batch on Multiple Device

Setup Alarm

Let us first get acquainted with the various terminologies used in setting up the Alarm.

Various Definitions:

| Alarm Type | To select the type of alarm that is either "High" or "Low". |
| --- | --- |
| Set Point | To set the set Point value for the alarm. The value to be considered to take action to raise the alarm or not. |
| Event | To Select the type of event that is either select S(Single) or select C(Cumulative). |
| Alarm Delay(hr:mm) | To Set the alarm delay time for the alarm. Not to take any action till the given time, if found any alarm set. |

Let us now configure the Alarm.

When the LM is online, (This feature is not available offline)

Click on LittleMaster  Setup Multiple LittleMaster

OR Click on the Icon  to setup batch on LittleMaster.

Use the Setup Alarm command to set the various parameters of the Batch.

if it is a "Temperature" Device then the following parameters need to be configured.

Like: Temperature

1. Alarm 1(A1)

2. Alarm 2(A2)

3. Alarm 3(A3)

4. Alarm 4(A4)

Particular parameters you can set in each alarm is as follows:

1. Alarm type

2. Set Point

3. Event

4. Alarm Delay(hh::mm)

Click on the Edit button to make all the parameters editable.

if it is a "Temperature & Humidity" Device then the following parameters need to be configured.

Like: Temperature

1. Alarm 1(A1)

2. Alarm 2(A2)

3. Alarm 3(A3)

4. Alarm 4(A4)

Humidity

1. Alarm 1(A5)

2. Alarm 2(A6)

3. Alarm 3(A7)

4. Alarm 4(A8)

Particular parameters you can set in each alarm is as follows:

1. Alarm type

2. Set Point

3. Event

4. Alarm Delay(hh::mm)

Click on the Edit button to make all the parameters editable.

If there is no batch running on LM then as a matter of caution following message is displayed.

This is the last chance for the user to download the previously stored batch. Once the new batch is configured, the data of the old batch will instantly lose. Select appropriate action to proceed.

If there is a batch running on LM then following message is displayed.

This message gives options to the user like to stop batch or stop and download batch before configuring new batch or cancel (Do not configure new batch).

- In edit mode, you can configure alarm data for four different alarm that is A1, A2, A3 and A4.

- User can edit Alarm type that is either "High" or "Low".

- Type of event you want to configure that is either "Single" or "Cumulative".

- Set point for the alarm and also the Alarm Delay time to be considered to trigger the alarm.

- After the configuration of alarm data click on " Configure Batch" and Setup Batch prompt will open to configure the batch data and to start the batch on the multiple devices.

- Click on the "Exit" button to close the Setup alarm prompt.

Setup Batch

Let us first get acquainted with the various terminologies used in setting up the batch.

Various Definitions:

| ID Name | Give the appropriate ID Name. Maximum 10 characters. It is generally the name which provide identification of device. Or it can be also the name of the location of the LM. It also appears in print report. |
| --- | --- |
| Batch Name | Give an appropriate name for the batch name. This is not reported in print reports. |
| Units of Temperature | Select appropriate Unit Temperature (Deg C/Deg F). |
| Store Interval | Time interval between two successive stored data in hr:mn:sc. Minimum store interval is 5 seconds and maximum is 18hrs. |
| No of Readings | These are the no of data sets that will be stored for the given configured batch. The maximum no of data sets is 50000* for temperature and %RH logger and 50000* for only temperature logger. This number is calculated based on the start, stop and store interval time. |
| Batch Type | User select whether to stop batch on memory full or Rollover data. |
| Time zone | Using time zone user can set the Current Time, Start Time and Stop Time according to different areas of the world. |
| UTC(Universal Time Coordinated) | The field is for the reference to set the current time accordingly. |
| Current Time | It is the current time of the LM device. If you want to set current time different, you can set it here. |
| Start Time | It is the time at which LM will start the batch. The first data is stored at start time of batch. |
| Stop Time | It is the time at which LM will stop storing data. |
| Display Auto Off | Select for auto off display of Device after batch config. It saves battery life of device. |
| Display Permanent On | Select for permanent on display of device after batch config. It decreases battery life of device. |

* Actual no. of data that can be stored is based on the product that you have purchased.

Multiple device batch configuration command uses for config same batch on multiple devices at a time.

When the LM is online, (This feature is not available offline)

Click on LittleMaster Setup Multiple LittleMaster Setup Alarm Configure Batch

Or click on  Setup Batch on Multiple Device button.

After the Setup Alarm prompt get open click on the edit button to allow user to edit parameters and enable Configure Batch button.

If there is no batch running on LM then as a matter of caution following message is displayed. if you want to download the batch data can go with yes button, otherwise click on No button and move forward.

After the configuration of the alarm data. Click on the "Configure Batch" button to open the "Setup Bach" prompt.

Setup Batch Prompt will open to configure the batch data.

- In edit mode user can set ID Number, Batch Name and Units for temperature and User can set the different time zone accordingly.

- The Current time shows the current time as on the LM. If user wants to change it, simply click on the clock icon next to it and it set the values of current RTC of PC.

- Also set the Start and Stop time for the Batch. Based on the start, stop time and the store interval, the no. of readings that will be stored in LM is displayed. If the no. of readings becomes more than the device capability, it will give an error message asking to modify any of the three parameters so that total no of readings is within limits.

- Select display option for device display auto off or permanent on.

- Once you are satisfied with the set parameters click on "Configure Batch" to set these parameters on LM.

- icon resets Current time and synchronizes with users PC time.

- Click on the Icon  to open calendar through which date and time can be set for Current Time, Start Time and Stop Time.

Note: For setup same batch on multiple devices, it's required that all device has same Sensor Type and Max no of Memory. In multiple device configuration process first device configuration considered as reference configuration.

- On click to "Setup Batch on Multiple Device" icon it will open Alarm Configuration dialog after that you need to click on Configure Batch for the configuration of batch.

- On click Configure Batch button it starts configuration process it shows you clear memory process information dialog. After that batch configuration process start.

- On Successful Configuration of the Batch, user will get the following message.

- On clicking ok it will show you the below “Multi Batch Config” dialog to select another connected device to setup same batch only if multiple LM devices are connected.

- Click on any connected device in which you want to config same batch as previous device and then click on Select to setup same batch configuration.

- On selecting the device application will show Setup Batch dialog to change into another suitable ID Number or Batch Name for selected device but all the others fields are gone be disabled as we are setting up same batch for all the connected devices. Then after click on “Configure Batch”.

- After all the connected devices are configured successfully application will show message box “All devices Configured Successfully”.

Note: Based on the Stop time, start time and Store interval, the no of readings are calculated. When the device is connected, this panel shows the time and parameters of the LMDevice. Please ensure the correct Start, stop time and no of readings are as per your desired settings. Click on "Configure Batch" only after you have verified all the parameters.

Batch Configuration for LM E051 Version-2 Devices:

- Users get an extra option for stop on full and roll over in Batch Configuration

- Select Batch Type "Rollover" or "Stop on Full":

- Rollover: When the maximum no. of readings is reached, data is overwritten cyclically on a first-in-first-out basis.

- Stop On Full: The batch stops when the maximum no. of readings is reached.

Note: If User select Batch Type as "Rollover", then stop time selection is disabled

Batch Configuration for LM-U Pro Devices:

Let us first get acquainted with the various terminologies used in setting up the batch.

Various Definitions:

| ID Name | Give the appropriate ID Name. Maximum 10 characters. It is generally the name which provide identification of device. Or it can be also the name of the location of the LM. It also appears in print reports. |
| --- | --- |
| Batch Name | Give an appropriate name for the batch name. This is not reported in print reports. |
| Alarm Temperature High | Set the required Alarm high limit here. When the temperature data crosses this limit, the display starts flashing. |
| Alarm Temperature Low | Set the required Alarm low limit here. When the temperature data goes below this limit, the display starts flashing. |
| Units of Temperature | You can select between °C or °F unit of measurement for temperature. Once the file is downloaded, user can convert to any unit, either °C; °K or °F on a click of a button. |
| Sensor Type | Select the type of senor attached to device. |
| Store Interval | Time interval between two successive stored data in hr:mn:sc. Minimum store interval is 10 seconds and maximum is 18hrs. |
| No of Readings | These are the no of data sets that will be stored for the given configured batch. The maximum no of data sets is 16000* for temperature and %RH logger and 32000* for only temperature logger. This number is calculated based on the start, stop and store interval time. |
| Current Time | It is the current time of the LM device. If you want to set current time different, you can set it here. |
| Start Time | It is the time at which LM will start the batch. The first data is stored at start time of batch. |
| Stop Time | It is the time at which LM will stop storing data. |
| Display Auto Off | Select for auto off display of Device after batch config. It saves battery life of device. |
| Display Permanent On | Select for permanent on display of device after batch config. It decreases battery life of device. |

* Actual no of data that can be stored is based on the product that you have purchased.

In edit mode user can set ID Name, Batch Name Alarm High and Low limits, RH High and Low limits and Units for temperature. The Current time shows the current time as on the LM. If user wants to change it, simply click on the clock icon next to it and if set the values of current RTC of PC. Also set the Start and Stop time for the Batch. Based on the start, stop time and the store interval, the no of readings that will be stored in LM is displayed.

If the no of readings becomes more than the device capability, it will give an error message asking to modify any of the three parameters so that total no of readings is within limits. Select display option for device display auto off or permanent on. Once you are satisfied with the set parameters click on "Configure Batch" to set these parameters on LM.

icon resets Current time and synchronizes with users PC time.

Click on the Icon  to open calendar through which date and time can be set for Current Time, Start Time and Stop Time.

# Download Batch from LM

Download the stored batch by clicking the Download stored batch icon. OR click on "Download Batch From LM" menu from menu bar as shown below.

Note: There are two type of download processes one is single file download and other is multiple file download. For Single file download it is required that the selected device should have same batch configuration with same Start Time and Store Intervals to download in single file.

Select the devices from the connected device to download the data from the device and then click on “Download”.

A Dialog box will open having the message “Do you want to save All the selected device data in SINGLE lm file”, On Clicking “Yes” all the selected devices data will be downloaded in single file.

The Save Combined LM File will open select the file location where you want to save the device batch data and enter the file name and then click on “Save”.

If any of the selected device for download will have different start time and store interval and the batch configuration does not match then above message will be displayed. On Click on “OK” the separate device lm data files will the download instead of single lm file. Or user can click on “Cancel” to stop download process.

On clicking “OK” the above window will display to select the folder to download the device data in separate lm files, click on “Select Folder” and then the download process will start.

The Download progress will be shown in download window and on successful download the success message will be displayed.

If the user Click on “NO” then instead of single device data file download system will download the seperated multiple device data files for each download selected devices.

Click on “NO” to Download the device data in multiple separate lm data files.

Select the folder to download the device data file in that particular folder location and click on Select Folder.

The downloading process will begin and after successful download the success message will be displayed on screen.

The lm files will be downloaded in LM_(SerialNumber)_(DateTime) format as you can see on above picture.

LmTransit Download Batch:

This feature is only for LmTransit Device.

Once the LMTransit device is found, the LmView automatically goes into download batch sequence. Following screen will come to save downloaded file. Select the appropriate location and name of the file to save the data.

After clicking on save button following screen will appear to configure the Batch on LM. Enter the required field.

User can only configure ID Name, Batch Name, Alarm Temperature High, Alarm Low and Timezone. All other parameters are preconfigured and are displayed to the user for his review.  Click on Save to download the data, data will download with selected timezone format. User can download same batch with different timezones for track data time on differet country. Ex. If one batch download by using timezone "(UTC+05:30) Chennai, Kolkata, Mumbai", it's batch data start from "13/06/2018 17:47:10" and same batch download by using timezone "(UTC-08:00) Pacific Time (US & Canada)", it's data start from "13/06/2018 05:17:10" again same batch download by using timezone "(UTC+10:00) Canberra, Melbourn, Sydny", it's data start from "13/06/2018 22:17:10".

After the download is over it will give a message "Download Successful". Click on "OK" to see the downloaded file data in tabular form.

In download data number of readings count depend on current time of device and start time of batch and it depends when you download data. On download if device current time greater than batch stop time then no of reading will be batch no of readings (Max memory of device) which was preconfigured. On download if device current time less then batch stops time then no of readings will be calculate as below. No of Readings = ((current time - start time) / Interval) + 1

# Calibrate LittleMaster

To prevent accidental change of the calibration, the calibration procedure is protected by a password. The default password is: gtek

Click on "Calibrate LM" menu from menu bar to open the Calibrate LM Dialog.

On click on the "Calibrate LM", Calibration Password dialog will appear.

click OK to enter calibration mode.

Temperature calibration:

- Select the Temperature in the choose sensor drop list.

- Enter the two reference values into the Reference value edit boxes.

- Enter the corresponding two LM values into the LittleMaster value edit boxes.

- Click Calibrate.

Humidity calibration:

To calibrate the humidity sensor, you will need a humidity chamber.

- Set the humidity chamber to the first reference value. Wait until the humidity level is stabilized and write down LM’s reading.

- Repeat (1) with the second reference value.

- Connect LM to the PC.

- Enter the calibration password, and then click OK.

- Select Humidity in the Choose sensor drop list

- Enter the two reference values into the Reference value edit boxes.

- Enter the two corresponding LM values into the LittleMaster value edit boxes.

- Click Calibrate.

Restore Factory Settings:

After giving the correct password click, Restore Factory Settings. This will reset the calibration parameter to original factory settings.

# Stop Running Batch

To stop the running batch, click onstop batch button. OR click the following menu from menu bar.

On applying stop batch command, it will show message which asks for conformation to stop running batch. On clicking "Yes" will stop batch and on clicking "No" will continue with batch storing.

After the batch is stopped successfully, following message will be displayed.

# Current Data

If the device is online, user can view the current temperature or temperature and %RH

(based on the device configuration) of LM by clicking on  icon in the toolbox.

User can view

- Current temperature reading.

- Current humidity reading, if LittleMaster has humidity sensor.

- Battery status of the battery of the LittleMaster.

Device Information like Product Code, Serial Number and Version.

User can switch between connected device to view the current data reading of that particular device.

The indicator shows the temperature and/or %RH. The units are as per the last configured batch.

Note: The "Product code" will be different according to the device. It will not be the same.

Indicator has feature to change its minimum and maximum value by right clicking on Indicator as shown below.

And

LM-U Pro Device:

The current data panel for the LM-U Pro is a bit different from other devices. The current data is displayed in Panel rather than a round dial.

User can switch between connected devices from the drop down to view the current data reading for that particular device.

The Panel have current date and time, alarm, live data reading, battery indicator, sensor type, temperature value as shown in below image

Note: User can see the over and under value in panel if the senor is not connected properly.

# Open Downloaded Batches

To show Downloaded Batch data into tabular form click on  Open Downloaded Batches icon. OR click the following menu from menu bar.

Select LM file which you want to open.

After selecting file click Open to see data into tabular form on main window.

Tabular View:

After the file is opened, the tabular data will appear on the right panel with the Minimum, Maximum, Average and MKT values, and on the left panel device information will be displayed.

Downloaded file contains either single device data like above image or multiple device data like in below image. In left panel it displays all device information and on right side it displays device data.

Tabular View with Tagged Data:

Tagged data display into tabular form with tagged data and the particular time and data.

Multiple device tag data like below the data row with the blue colour is the tagged data.

Change Units:

Once the tabular data is displayed, right click on the mouse button and Select option "Select Units".

if temperature device, select the desired unit. The columns will be converted along with minimum and maximum value for that data set.

If the %RH is there, you can also convert the columns of dew point. For calculation of the dew point, the corresponding temperature data of that record is used.

# Filter Table Data

Let us say you want to view / plot only selected part of the data. Right click on the tabular data or click on  filter icon to open filter window.

Filter Tabular Data option is enabled when data is displayed in tabular form in main window. For filter data required parameters are start date and end date. If user enters the start date and/or end date which is beyond the range of original data, the error message will be generated.

In the above batch there are 8497 readings with start time as 23-12-2022 10:45:00 and stop time as 24-12-2022 10:21:30.

for filter

Click "filter" command and let us change Start Time and End Time as indicated above and press "OK". The tabular data will be seen as under according to data filtered from start time "23-12-2022 10:45:00 " to end time "23-12-2022 10:46:30" in the below image.

It is important to know that the data is not deleted from the file. Whenever you open the file again, the total data will be displayed.

Note: if user have filtered the data to view, it will not affect on the alarm report generation.

# Format Tabular Data

To Let us say you want to format DateTime or Data decimal value in tabular form data, click the following menu from menu bar.

Data Formats menu is enabled when data is displayed in tabular form in main window. Using Data Format user can update DateTime and Data decimal values in tabular data. On click on Data Formats menu it will open the following dialog.

User can able to update datetime and data format using above dialog. Available DateTime and Decimal format as shown below.

Open Data Formats dialog and select any Date Format from drop down and press "OK".

The tabular data will be displayed as in the particular format selected by the user on press OK in Data Format dialog.

Open Data Formats dialog and select any Decimal value from drop down and press "OK".

The tabular data will be displayed in the particular decimal value form on press OK in Data Format dialog.

It is important to know that the selected datetime and data decimal format apply on tabular data and save. Whenever you open any file data in tabular form will be display in the same format selected by the user. Data decimal format will be applied on all type of reports.

# Calculate MKT

To calculate the MKT, click on  MKT icon, and enter the desired activation energy. The MKT is calculated and displayed in the box. It is also exported in .csv format and all other report.

After applying MKT command, shows following result.

# Unit Conversion

Click on  to convert the temperature from one unit to another.

OR click on the "Converts Units" menu from menu bar.

In above dialog, input temperature value to convert from input box and its unit. Select the desired unit you want to convert to. The converted temperature will be displayed.

# View Graph

Click on the icon  and the graph for all the displayed columns in table will be displayed in a new window. Alternatively, Graph can also be opened from the menu by right clicking on tabular data.

Graph window appears as shown.

| Part | Description |
| --- | --- |
| 1. Menu bar | The area containing pull-down menu options. |
| 2. Toolbar | A detachable bar that contains shortcuts to menu and other commands. |
| 3. Chart | The chart area to view data in graphical format. |

Display batch data and tagged data into graphical view as like above.

# Zoom In Graph

You can change the view of a graph by zooming in to get a closer look. You can experiment with a variety of Zoom In options like Zoom in Both Axis, X Axis and Y Axis to determine the amount of detail you want.

From the Zoom Option menu click on the   icon will zoom in graph from its center position or drag the cursor diagonally from left to right to select the area you want to magnify. Release the mouse button to zoom in both axis to the selected area.

X Axis:

On the Zoom Option menu click on Zoom In >   icon will zoom in X Axis of graph from its centre position.

Y Axis:

On the Zoom Option menu click on Zoom In >   icon will zoom in Y Axis of graph from its centre position.

# Zoom Out Graph

You can change the view of a graph by zooming out to get a previous view. You can experiment with a variety of Zoom Out options like Zoom Out Both Axis, X Axis and Y Axis to determine the amount of detail you want.

Both Axis:

On the Zoom Option menu click on   icon will zoom out graph and show its Both Axis on its original position.

X Axis:

On the Zoom Option menu click on Zoom Out >  will zoom out X Axis of graph and set X Axis to its original position.

Y Axis:

On the Zoom Option menu click on Zoom Out >  will zoom out Y Axis and set Y Axis of its original position.

# Save Graph

From the Graph utilities menu bar, click on save  icon

or right click on graph menu, click on Save Image As

It will open following Save graph window. You can give any file name and also select type of file(.emf/.png/.gif/.jpg/.tif/.bmp) and after clicking on Save button, graph image will be saved at your desired location and by the given name.

# Print Graph

From the Graph utilities menu bar, click on Print  icon.

or right click on graph menu, click on print.

It will open following Print window.

The default printer will be shown. User can select any of the installed printer from the dropdown menu. Click on OK button and current graph image will be printed on the selected printer. User can also change the properties of the printer by clicking on properties

# Transfer Axis

Lot of time it is useful to have the secondary Y axis for some of the data. Transfer Axis exactly allows you to do that. User can set as many channels as he desires for Y2 axis.

From the Graph utilities menu bar, click on Transfer Axis icon , the following Transfer Axis window will open. From here user can select either Y Axis or Y2 Axis for any channel and press on OK button.

The channel that is transferred to Y2 axis will be indicated in label as above.

# Select Channels

To remove the channel or add back the channel from the graph window, use Select Channels Menu.

From the Graph utilities menu bar, click on Select Channels icon ,the following Select Channels window will open. From here user can select channels for On/Off purpose.

# Graph Properties

You can set various graph properties like, Title of the graph, The Axis names, Axis scales and line colours. Following window shows details of chart elements which can be modified by Graph Properties dialog.

On the Graph utilities menu click on Graph Properties icon  , and it will open Graph Properties window.

Editing Chart Tittle:

A main chart tittle that can be changed by following chart properties dialog. You can also change its font and font colour.

Select the Title Tab, and check show title. In the Text field, give the desired title. Click on "Select" button to change the font size and colour.

Editing Axis:

Select the "Plot" tab, and on the XY plot select either Domain Axis (Time) or Range Axis Y (Temperature) or Range Axis Y2 () to set various parameters. Following window shows which properties can be set for the Axis.

Select Plot Domain Axis will change X axis tittle, its font style and font colour. Tick properties like visible or not and the font style can be set here.

Select Plot  Range Y will change Y axis tittle, its font style and font color. Tick properties like visible or not and the font style can be set here.

Select Plot Range Y2 will change Y2 axis tittle, its font style and font color. Tick properties like visible or not and the font style can be set here.

Click on Range Y2 Axis. In the Range, Deselect Auto adjust Range, Give the desired minimum and maximum range values in the edit box.

Changing Scale of Range Axis:

Only Range Axis Y and Range Axis Y2 have the scale properties.

Click on Range Y Axis. In the Range, Deselect Auto adjust Range, Give the desired minimum and maximum range values in the edit box.

Click on Range Y2 Axis. In the Range, Deselect Auto adjust Range, Give the desired minimum and maximum range values in the edit box.

Chart Appearance:

Select the "Plot" tab, and on the XY plot select Appearance to set Chart, ChartArea and Chart Series display properties. Following window shows various properties that can be set.

Select Appearance Chart will change Chart back color, border line style, line color and line width. After setting all these properties click on ok button and chart appearance is updated. Experiment with various settings to get the desired result.

Select Appearance ChartArea will change ChartArea back color, border line style, line color and line width. After setting all these properties click on ok button and chart appearance is updated. Experiment with various settings to get the desired result.

Select Appearance Series will open color dialog to change its color properties. After setting all these properties click on ok button and chart appearance is updated. Experiment with various settings to get the desired result.

# Export to CSV

Data can be exported in CSV format by either clicking  on the main toolbar or by File Export menu or by clicking Export to CSV popup menu on right click of tabular data windows.

It will open MKT calculation window.

Give desired activation energy and click "OK". MKT value in tabular data will be updated. Save file dialog will open, for saving .csv file on desired path.

After clicking on Save, displayed tabular data of the file will be exported in .csv format to the given file name and path, and following message will appear.

# PDF Report

Various Definitions:

| Report Title | Give an appropriate ReportTitle for the batch name. Maximum 50 character. |
| --- | --- |
| Company Name | Give the appropriate Company Name. Maximum 25 characters. This will be the title on the first page and footer in subsequent pages. |
| Additional Information | Give Summary report of data. Maximum 1100 character. |
| Activation Energy | Give activation energy value for calculate MKT. |
| Select Report Type | Select reports for create in current report. |
| Alarm Included in Report | Select the number of alarms information to include or to skip the information of the particular alarm to not to be included while generating report. |
| Observation | Input any observation regarding to report. |
| Conclusion | Input any conclusion regarding report. |
| Generate Report | Click on Report to generate PDF report. |
| Header Footer | Click on Header Footer for insert custom header footer in report. PDF report. |
| Cancel | Click on Cancel to exit form Report window. |

Data can be exported in PDF format by either clicking   on the main toolbar or by right click on tabular data open menu.

After clicking on save button it will open following Generate PDF Report window.

For Temperature and Humidity device.

- You can set Report Title, Company Name and Additional Information. Also set activation energy if create MKT report.

- Select required reports from Report list, and also You can Select the Alarms to be included in the Report accordingly so only that alarm data will be included in all the reports you will be selecting, rest of the alarms data will not be included in the report.

- Once you are satisfied with the set parameters click on "Generate Report" to generate pdf report.

Recommendation: Whenever user Selects the Alarm to be included in the report if user Select Alarm1 do select Alarm5 as well, as shown in the above image in the same way if you select Alarm2 do Select Alarm6 as well and so on for the better and proper information while generating the "Report".

Explanation and Example of Types of Report:

- Summary Report

- Reports contains the minimum and maximum temperature of the data along with the recorded time, also contains the average temperature, Mean Std temperature and MKT temperature as shown below in the image.

- Also contains the above-mentioned information for the Humidity as shown below in the image

- Alarm Report

- Alarm Report will contain the details of the alarm configuration that is the type of alarm and the type of event.

- Also, the Total time of the alarm continued to trigger, first triggered time of the alarm, no of occurrences of the alarm and lastly the status of the alarm if alarm is there it will show the status "Alarm", if not the status will be "ok".

- Alarm Report will only contain the information of the alarms included at the time of report generation.

- Data Report

- Data Report contains the Batch data in the tabular form.

- Data Graph

- Displays the Batch data in the Graphical view as shown in the below image.

- Tag Data

- It contains the information about all the tags and the data from batch data along with the particular time.

- Also contains the description of each tag for understanding of the user.

- Tag Data Graph

For Temperature device

Following screen will come to save pdf file. Select the appropriate location and name of the file to save the data.

On Successful generation of pdf report, user will get the following message.

The displayed tabular data file will be exported in .pdf format at given path and name.

Header Footer:

- In Header Footer dialog it provides selection for header and footer to set it as customize or fixed. There are two different sections for set Header and footer.

- In Header section there are two option available one is Fixed Header and other is Image Header.

- On selection of Fixed Header, it will create report using our fixed format of Header.

- On selection of Image Header, it will enable Load Image button which provide image selection dialog. On selection of any image it will show into image preview area like as below image and it will create report using this image as Header.

- In Footer section there are two option available one is Fixed Footer and other is Custom Footer.

- On selection of Fixed Header, it will create report using our fixed format of Footer.

- On selection of Custom Footer, it will enable text area to write any content which use as footer. Any inserted content into text area view as footer in report page.

# Error Description

Error Communicating with Device is a generic error. This error can occur at various stage of communicating to device. The following table indicates the place of error occurring.

| Error Code | Error Message | Error Description |
| --- | --- | --- |
| C0001 | Error Communicating with Device.... Switching to offline mode | This error is generated when device had communication problem during "setting Calibration value" |
| C0003 | Error Communicating with Device.... Switching to offline mode | This error is generated when device had communication problem during "Restore factory settings" |
| C0004 | Error Communicating with Device.... Switching to offline mode | This error is generated when device had communication problem during "Setup Batch Configuration" LMTransit device. |
| C0005 | Error Communicating with Device.... Switching to offline mode | This error is generated when device had communication problem during "download Stored Batch" from LittleMaster device. |
| C0006 | Error Communicating with Device.... Switching to offline mode | This error is generated when device had communication problem during "show the Current Data" |
| C0007 | Error Communicating with Device.... Switching to offline mode | This error is generated when device had communication problem during "Setup Batch Configuration" on LM device |
| C0008 | Error Communicating with Device.... Switching to offline mode | This error is generated when device had communication problem during "Stop Batch" command on LM device. |
| C0009 | Error Communicating with Device.... Switching to offline mode | This error is generated when device had communication problem during "Waking up the device" and application goes into offline mode. |

Remedies:

Whenever there is an error communication message, please check that the communication cable is properly connected to both the PC and LM end. In case of LmTransit please ensure that the jack is fully inserted in. After verifying it, please try again to wake up the device. If the problem still persists, remove the battery and re-insert or replace in case the battery is old. And try again.

The device automatically goes in to battery saving mode if no activity takes place for more than 5 minutes after waking it up. In this case also the "Error communicating message will come. To resolve this, "go to offline mode" and again go to "Online mode" this will wake up the device again and communication shall resume.

(If the problem still persists, please send the error message no with the log file (..\ProgramFiles\LMView\LMView.bin) file to

service@gtek-india.com. We shall try to respond you within 24 hours.